## Slide 1: SuperBigBite DAQ update

## Slide 2: Outline
SBS requirements
Fastbus Readout
GEM readout
FADC HCAL readout

## Slide 3: SBS DAQ requirements
Gep5
  200 KHz L1 Electromagnetic Calorimeter Fastbus
  Coordinate detector readout
  GEM readout VME
  HCAL trigger pipeline FADC and L2 trigger
GEn / GMp
  Bigbite readout Fastbus
  HCAL as neutron detector using FADC and high resolution TDC

## Slide 4: Fastbus update
Use new CODA 3.0 TI and TS
  More flexibility in programming
  Event blocking
  Absolute timestamp for synchronization check
  Trigger partitioning capability only use a subset of modules
Asked for modified firmware to DAQ group in May 2015 ( William Gu and Bryan Moffit )
Firmware being developped and tested

## Slide 5: Fastbus status @ JLab
Have sufficient TDCs, ADCs, crates, aux. cards
Like to scrounge ~3 more SFI (Struck Fastbus Interface)
(backup plan: fewer crates, reduced performance)
Making FB faster
Three large Fastbus systems are being
assembled for test in the test lab.
sparsification – works
event blocking – works
event switching – being tested
merging with pipelining VME -- to be tried
have 113
need 94
have 236
need 124
have 30
need 21

## Slide 6: GEp DAQ Configuration / both arms
T2 distributed to Hadron and Electron Arms;
if no L2A, Fast Clear to Fastbus
L2
L2
L2
Fastbus

## Slide 8: Single Crate, Buffer Level = 4
Chart: Readout trigger rate (Hz)
| 1145.4000000000001 | 8119.1 | 10814.7 | 4913.8999999999996 | 13093.1 | 16890 | 22452.5
Local_50kHz_pedeSup | 10.11 | 14.5977 | 15.7434 | 12.9205 | 17.3198 | 23.0432 | 34.8999
Local_100kHz_pedeSup | 18.5354 | 18.9397 | 20.5634 | 21.208 | 23.132 | 28.9388 | 37.0528
Local_50kHz_pedeSup | 25.6377 | 26.7435 | 27.9307 | 28.9473 | 30.2244 | 35.8366 | 40.8775
Local_50kHz_NopedSup | 24.8035 | 42.9484 | 12.4653 | 10.0686 | 57.4176 | 72.4219 | 81.83 | 85.1154 | 13.2788 | 12.0928
Local_100kHz_NopedSup | 19.7471 | 20.6118 | 22.0903 | 37.7701 | 53.8329 | 67.5743 | 78.1296 | 84.5678
Local_200kHz_NopedSup | 25.1119 | 26.1958 | 38.9807 | 52.4029 | 34.3715 | 69.9737 | 79.109 | 28.0518
No Pedestal
Suppression (8ADCs)
Pedestal
Suppressed
Local trigger ~ 100 kHz
Local trigger ~ 50 kHz
Local trigger ~ 200 kHz
Expected ~ 5000 Hz
Dead time at
Local ~ 200 kHz and Readout ~5kHz
D1 ~40%

## Slide 9
3 Crates, Buffer Level = 4
Chart: Readout trigger rate (Hz)
| 1179.4000000000001 | 1302.5999999999999 | 2259.6999999999998 | 3761.9 | 3194 | 7769.2 | 12436.7 | 17482.099999999999
Local_50kHz_pedSup | 0.9242 | 1.1515 | 1.9604 | 2.1213 | 2.0382 | 2.3336 | 4.0075 | 11.6422
Local_100kHz_pedSup | 3.2204 | 3.8407 | 4.1558 | 4.4367 | 5.2404 | 12.845 | 0 | 0 | 0
Local_200kHz_pedSup | 5.8897 | 6.0542 | 5.6354 | 5.4197 | 6.5448 | 10.4784
Local_50kHz_NopedSup | 2.22 | 1.4363 | 1.0648 | 31.7394 | 15.4081 | 1.3019 | 3.5282 | 32.2435 | 4.468 | 38.856 | 0 | 0.73978 | 8.6514 | 20.8324 | 7.9534
Local_50kHz_NopedSup | 2.338 | 2.2167 | 23.8158 | 9.2238 | 37.8864 | 44.0168 | 34.4805 | 31.8351 | 8.8881 | 6.7969 | 4.1572 | 2.8746 | 1.4263
Local_200kHz_NopedSup | 4.5394 | 5.1787 | 0 | 5.7375 | 10.9662 | 20.8511 | 49.0975 | 31.8612 | 12.6461 | 56.7151 | 63.4492 | 7.5695
No Pedestal
Suppression (8ADCs)
Pedestal
Suppressed
Local trigger ~ 100 kHz
Local trigger ~ 50 kHz
Local trigger ~ 200 kHz
Expected ~ 5000 Hz
Dead time at
Local ~ 200 kHz and Readout ~5kHz
D3 ~6%
D3 ~ (D1)3

## Slide 10: Readout trigger rate ~ 5kHz, Buffer Level = 4, No pedestal suppression (8ADC)
Chart: Local trigger rate (Hz)
| 44669.2 | 104913.7 | 150700.4 | 194717.3 | 27297.9 | 11369.4 | 199979.1 | 170920.4 | 65165.5 | 83922.7 | 8182.8
buffer4_3daq_nopedsup | 2.5099 | 3.7419 | 4.1961 | 6.5882 | 2.0287 | 1.3511 | 6.8404 | 5.2867 | 3.0341 | 3.6757 | 1.1329
 | 1.7391 | 2.3606 | 2.6502 | 4.4902 | 6.5756 | 9.4732 | 11.1386 | 17.5771
 | 26.3999 | 26.1214 | 27.5933 | 0 | 28.7049 | 28.0324 | 28.1751
3 crates
2 crates
1 crates

## Slide 11: HCAL
All FADC 16 delivered
2 VXS crates delivered
2 Intel Concurrent CPU delivered
Readout tested
Development of trigger using HPS firmware and GTP
New VTP ordered

## Slide 12: HCAL FADC electronics
Cosmics from calorimeter block
Will test later with HCAL module

## Slide 13: Hadron Arm - HCAL DAQ: proton trigger
2 VME switched Serial (VXS) Crates
JLAB FADC250, a 16-channel 12-bit FADC sampling at 250 MHz
  16 FADC in VXS Crate 1
  2 FADC in VXS Crate 2
  If signal pass threshold
    Integrates signal and subtracts pedestal
    Sends time frame info
HCal Signals to FADC inputs
VTP (Crate Trigger Processor)
  Located in VXS crates
  Collects integrated signal and timing from FADC channels
  Sends data to the second VTP for clustering

## Slide 14: HCAL trigger development
HPS firmware installed on FADC and GTP

## Slide 15: GEM MPD readout
INFN MPD used for several years using custom C++ package
Package ported to intel CPU
New C library written for easy integration into CODA
( Bryan Moffit, Evarsito Cisbani, Danning Di )
CODA configuration running
Debugging module initialization

## Slide 16: Status MPD
VME part seems to work
Maybe issue with communication with APV being investigated

## Slide 17: GEM optical link readout
Aurora protocol based
SSP readout library being tested
Link from MPD to SSP module working from Ben Raydo
MPD functions implementation in Optical module in progress by Paolo Musico

## Slide 18: Timeline

## Slide 20: Conclusion
Fastbus flipping module close to be complete
  Preliminary results show acceptable dead time in experiment conditions
  Need to develop software and check synchronization
HCAL : FADC ready, trigger implemented and being tested
MPD : CODA readout implemented, debugging of the C code, Optical readout in progress
Small scale test setup in a few weeks